## Slide 1: Towards a generic dispenser module for laboratory platforms
Lars Yndal Sørensen, ITU

## Slide 2: This presentation…
Design goals
  Based on currently available solutions
Proposed solution
  Demo! :-D
Core parts of the design
Evaluation
Future aspects

## Slide 3: Design goals
Modularity Attach/remove module without hassle.
Versatility Different sizes of Petri dishes/multiwell plates.
Usability Easy to add/remove stacks during runtime.
Flexibility Can be attached to most current laboratory platforms.
Increased feedback Detects platform layers automatically.
Repairability Easy to create spare parts - and low cost.
Responsive Control Responses from module, regardless of ongoing motions, etc.

## Slide 4: Proposed solution
Demo of:
  Calibration
  Detecting layer
  Detecting dish stack
  Defining insertion point
  Inserting dish

## Slide 5: Core parts – the hardware
Arduino Mega vs Teensy
  RAMPS vs DIY, speed, cost, size, no reset upon connection
Electrical circuits for stacks
Low use of screws: Self-locking design
Stacks
  Self-centering, automatic identification, easy access for user
Gripper
  Squeezing
Protects endstops from damage
95% open-source parts
  Parallax PING))) and microbarbs are not

## Slide 6: Core parts – the software
Space conserving
Implements acceleration/deceleration
Self-calibrating
Automatic stack detection
Responsive (but not when calibrating!)
Easy to add new motions
Provides useful feedback

## Slide 7: Evaluation
Fulfilled design goals:
  Modularity No hassle
  Versatility Dish sizes and types
  Usability Add/remove stacks during runtime
  Flexibility Fits most current laboratory platforms
  Increased feedback Detects platform layers automatically.
  Repairability Easy to create spare parts - and low cost.
  Responsive Control Responses from module
6/7 goals achieved!

## Slide 8: Issues to be addressed
  Angle of lead screws
  Support blocks for lead screw
  Lifting wells (squared objects)
  Risk of power drop, when inserting stacks
  Small stepper gets hot
  Organize the wires!!

## Slide 9: Possible improvements
Save calibration data in EEPROM
Concurrent motions
Install electronics more safely (wires + MC)
Ball bearings for supporting the leadscrews
How to grip bottom and top of dish in one go
Design that uses less resources

## Slide 10: Future aspects
Used by small to medium companies
  Small companies can afford it and it supports different types of sample containers (almost…)
  Extendable in size (and has space to add extra equipment)
Used by upcoming/current research projects
  Easy to setup for even a small batch of samples

## Slide 11: ERRATA
Hall effect sensors tested:
  3144 OH14 is the switch and and 49E 11082N is the linear